## Slide 1: Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
MEng Individual Project Presentation
Yifei Jin
A Promising Technique for Parkinson’s Disease Biomarker Detection

## Slide 2
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…
Notes: In this presentation, I will first introduce a bit of the background information, followed by brief explain of the theory of N-periodic ka-SPGR sequence T2* mapping. Then I will get into the method and result part, which spilt into computer simulation and practical MRI scan. After that, I will conclude the result of this project and also the limitations.

## Slide 3
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…
Notes: First, lets go into the background introduction.

## Slide 4
Background introduction
Parkinson’s Disease
Excess iron at Substantia Nigra (SN)
Quantitative R2* images of SN
Healthy control T2* | PD patient T2*
53 ms | 13 ms
Brammerloh, Malte et al.(2021) [1]
MRI time constant T2* as an effective biomarker
R2*
Notes: Parkinson’s Disease is second most common neural degenerative disease, which affect the motor region of brain. It is irreversible and currently has no cure. Therefore, early detection is important, as it will help treatment research and tracks disease development. At early onset of Parkinson’s Disease, increasing iron concentration is observed at the substantia nigra region in midbrain. It distort the magnetic field when doing MRI and can be detected using an MRI time constant T2*, which is tissue and field inhomogeneities dependent. Previous research had shown the T2* value drop from 53ms to 13ms for PD patient, and you can see these change quite obviously in the substantia nigra in these figure below, which are R2* quantitative MRI image, R2* is just reciprocal of T2*. So, this drop in T2* can be used as an effective biomarker for PD. But, what exactly is the T2*? And how do we calculate the T2* and generate the T2* map?

## Slide 5
Background introduction
Gold-standard T2* mapping technique
What is T2*?
Signal Intensity (SI)
Time
Multi-echo spoiled gradient recall
(multi-echo SPGR)
TE = a
TE = b
TE = a
TE = b
repeat…
Notes: To generate MR image, we excite the magnetised spins with certain flip angle using radiofrequency pulses, and record the signal at echo time TE in Fourier space called k-space during relaxation. And an inverse Fourier transform of k-space is performed to reconstruct the MR image. The recorded signal intensity exponentially decayed against TE with time constant T2*. By acquiring signal at multiple different TEs, we will be able to fit the exponential decay and calculate the T2* value voxel by voxel. And this is the gold-standard T2* mapping method called multi-echo SPGR.

## Slide 6
Background introduction
Gold-standard T2* mapping technique
What is T2*?
Signal Intensity (SI)
Time
Multi-echo spoiled gradient recall
(multi-echo SPGR)
TE = a
TE = b
TE = a
TE = b
repeat…
Motion artefact
Possible solution:
N-periodic ka-SPGR
Holdsworth, S. J et al. (2016) [2]
Notes: However, the gold-standard method is difficult to get high resolution and accurate T2* map at SN due to the tissue motion at brain stem you can see here. A method based on a more complexed N-periodic ka-SPGR sequence is proposed to be motion robust and can be used to get T2* map.

## Slide 7
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…
Notes: So, now just let me briefly introduce the ka-SPGR sequence.

## Slide 8
Theory
N-periodic ka-SPGR
S(0)
Small flip angle
Short repetition time TR
TE = TR/2
RF spoiling
…
…
S(1)
S(N-1)
S(0)
N-periodic k-space aliased SPGR (ka-SPGR)
0 (N-1)
- Gradient spoiling
Pete, Lally et al.(2022) [3]
Notes: For a N-periodic ka-SPGR sequence, we excite the spins with small flip angle, short TR and a TE equals half TR, also, a RF-spoiling is added, which means excitation each time will have phase increment according to this quadratic equation. Using all of these settings, the signal generated are no longer directly related to T2* decay. Instead, it yields N different steady-state signals from S(0) to S(N-1), and each steady-state signals are the RF phase modulation weighted sum of exponential T2* decay signal at specific times, shown in this analytical solution of N-periodic ka-SPGR. This yellow term, the exponential decay signal at TE + F TR are called configuration state, or F-states. The F-states can be reconstructed by correcting the phase modulation. However, using N-periodic ka-SPGR will only allow acquiring N-steady state signal, and is not able to reconstruct more than N F-states. Therefore, we need to shift the F-states using a gradient spoiler as shown here, to make sure only N F-states are inside the acquired k-space.

## Slide 9
Theory
N-periodic ka-SPGR
F = 0
Correct for
phase modulation
F = 1
F = 2
F = 3
F = 4
Shift
F-states
(image-space)
F0
F1
F2
F3
F4
n = 0
n = 1
n = 2
n = 3
n = 4
Reconstruction example:
5-Periodic ka-SPGR
F0
F1
F2
F3
F4
F-state magnitude
Time
Notes: An example of reconstructing the F-states and fitting the T2* value is shown here using 5-periodic ka-SPGR. We correct the phase modulation of each acquired signal using this equation and sum up to form the k-space of corresponding F-states. Then we shift the off-set F-states information back to the middle of k-space and inverse Fourier transform to get the F-states signal. The magnitude of F-states are then used to fit exponential decay against the effective time TE+FTR and acquire T2* as shown here.
As we construct the signal used to fit T2* decay is averaged across multiple acquisition, this method is theoretically motion robust. However, the accuracy of the T2* mapping of this technique is not quantitatively analysed and no optimal scanning parameters are suggested before.

## Slide 10
Method Overview
Monte Carlo computer simulation
Optimised TR and N for N-periodic ka-SPGR
Practical phantom scanning experiment
Percentage bias
Bias and variation
T2* measurement efficiency
T2* mapping image
Notes: In this project, we first build a computer simulation to calculate bias and find optimal scanning parameters at lowest bias. Then the optimized sequences are used to perform a practical phantom scanning experiment to generating T2* image and analyse the performance in reality comparing with gold-standard method.

## Slide 11
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…

## Slide 12
Computer Simulation
Method
T2* ground truth (13-53 ms)
TR
(2-10ms)
1000 measured T2*
Monte Carlo experiment data
T1 | 800 ms
T2 | 44 ms
Flip angle | 10
For each N in range 3-12

## Slide 13
Computer Simulation
Result
Average T2* percentage bias

## Slide 14
Computer Simulation
Result
Optimal ka-SPGR percentage bias and standard deviation
N = 12, TR = 6ms
N = 7, TR = 6ms

## Slide 15
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…

## Slide 16
Practical MRI Scan
Method
Multi-echo SPGR
N-periodic ka-SPGR
Time
SI
TE1
TE2
TE3
TE
TE + TR
TEN
TE + 2TR
TE + (N-1)TR
……
……
……

## Slide 17
Practical MRI Scan
Result
Gold-standard
Quantitative T2* mapping images
7-periodic ka-SPGR
12-periodic ka-SPGR
Notes: Percentage error, middle line zero bias, red line +-10%
12 periodic has smaller std
All within +-10% bias
Effective T2* SNR
 7 periodic more efficient at low T2*

## Slide 18
Practical MRI Scan
Result
T2* percentage error
N = 7
N = 12
Effective T2* SNR
Notes: Percentage error, middle line zero bias, red line +-10%
12 periodic has smaller std
All within +-10% bias
Effective T2* SNR
 7 periodic more efficient at low T2*

## Slide 19
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…

## Slide 20
Conclusion
Conclusion
Optimal ka-SPGR:
7-periodic ka-SPGR TR = 6ms
12-periodic ka-SPGR TR = 6ms
7-periodic is more efficient for PD detection.
12-periodic is more precise.

## Slide 21
Limitation and Further
Limitation
Phantom misalign with substantia nigra tissue property (T1).
Large amounts of in vivo scan experiments are required to statistically prove the method for clinical use.
Further
The optimal parameters can be used to test motion robustness in vivo.

## Slide 22: References:
Brammerloh, M., Morawski, M., Friedrich, I., Reinert, T., Lange, C., Pelicon, P., Vavpetič, P., Jankuhn, S., Jäger, C., Alkemade, A., Balesar, R., Pine, K., Gavriilidis, F., Trampel, R., Reimer, E., Arendt, T., Weiskopf, N., & Kirilina, E. (2021). Measuring the iron content of dopaminergic neurons in substantia nigra with MRI relaxometry. NeuroImage, 239. https://doi.org/10.1016/j.neuroimage.2021.118255
Holdsworth, S. J., Rahimi, M. S., Ni, W. W., Zaharchuk, G., & Moseley, M. E. (2016). Amplified magnetic resonance imaging (aMRI). Magnetic Resonance in Medicine, 75(6), 2245–2254. https://doi.org/10.1002/mrm.26142
Lally, P. J., Chiew, M., Matthews, P. M., Miller, K. L., & Bangerter, N. K. (n.d.). A motion-robust, short-TR alternative to multi-echo SPGR.
References