## Slide 1: Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
MEng Individual Project Presentation
Yifei Jin
A Promising Technique for Parkinson’s Disease Biomarker Detection

## Slide 2
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…
Notes: In this presentation, I will first introduce a bit of the background information, followed by brief explain of the theory of N-periodic ka-SPGR sequence T2* mapping. Then I will get into the method and result part, which spilt into computer simulation and practical MRI scan. After that, I will conclude the result of this project and also the limitations.

## Slide 3
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…
Notes: First, lets go into the background introduction.

## Slide 4
Background introduction
Parkinson’s Disease
Excess iron at Substantia Nigra (SN)
Quantitative R2* (1/T2*) images of SN
Healthy control T2* | PD patient T2*
53 ms | 13 ms
Brammerloh, Malte et al.(2021) [1]
MRI time constant T2* as an effective biomarker
R2* = 1/T2*
Notes: Parkinson’s Disease is the second most common neural degenerative disease, it is irreversible and currently has no cure. Therefore, finding an effective biomarker for early detection is important. At the early onset of Parkinson’s Disease, increasing iron concentration is observed at the substantia nigra region in the midbrain. It distorts the magnetic field when doing MRI and can be detected using an MRI time constant T2*, which tissue and field inhomogeneities dependent. Previous research had shown the T2* value drop from 53ms to 13ms for PD patients, and you can see here in these quantitative mapping images, the change is quite obvious. So, this drop in T2* can be used as an effective biomarker for PD. But, what exactly is the T2*? And how do we calculate the T2* and generate the T2* map?

## Slide 5
Background introduction
What is T2*?
Signal Intensity (SI)
Time
TE
TE
repeat…
TR
image-space
k-space
Notes: To generate MR image, we excite the magnetised spins with a certain flip angle using radiofrequency pulses and record the signal at echo time TE in Fourier space called k-space. And an inverse Fourier transform of k-space is performed to reconstruct the MR image. The recorded signal intensity exponentially decayed against TE with time constant T2*.

## Slide 6
Background introduction
Gold-standard T2* mapping technique
What is T2*?
Signal Intensity (SI)
Time
Multi-echo spoiled gradient recall
(multi-echo SPGR)
TE = a
TE = b
TE = a
TE = b
repeat…
Notes: By acquiring signal at multiple different TEs, we will be able to fit the exponential decay and calculate the T2* value voxel by voxel. And this is the gold-standard T2* mapping method called multi-echo SPGR.

## Slide 7
Background introduction
Gold-standard T2* mapping technique
What is T2*?
Signal Intensity (SI)
Time
Multi-echo spoiled gradient recall
(multi-echo SPGR)
TE = a
TE = b
TE = a
TE = b
repeat…
Motion artefact
Possible solution:
N-periodic ka-SPGR
Holdsworth, S. J et al. (2016) [2]
Notes: However, the gold-standard method is difficult to get high resolution and accurate T2* map at SN due to the tissue motion at brain stem you can see here. A method based on a more complexed N-periodic ka-SPGR sequence is proposed to be motion robust and can be used to get T2* map.

## Slide 8
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…
Notes: So, now just let me briefly introduce the ka-SPGR sequence.

## Slide 9
Theory
N-periodic ka-SPGR
Small flip angle
Short repetition time TR
TE = TR/2
RF spoiling
N-periodic k-space aliased SPGR (ka-SPGR)
S(0)
…
S(1)
S(N-1)
S(0)
…
Pete, Lally et al.(2022) [3]
…
Notes: For a N-periodic ka-SPGR sequence, we excite the spins with small flip angle, short TR and a TE equals half TR, also, a RF-spoiling is added, which means excitation each time will have phase increment according to this quadratic equation. Using all of these settings, the signal generated are no longer directly related to T2* decay. Instead, it yields N different steady-state signals from S(0) to S(N-1).

## Slide 10
Theory
N-periodic ka-SPGR
S(0)
…
S(1)
S(N-1)
S(0)
N-periodic k-space aliased SPGR (ka-SPGR)
0 (N-1)
Pete, Lally et al.(2022) [3]
…
…
Adjust gradient spoiling
Notes: Each steady-state signal has the analytical solution listed here. This yellow term is called configuration state, or F-states, it is the exponential T2* decay signal at effective time TE + FTR. Each steady-state signal is the sum of F-states weighted by corresponding phase modulation.
Using the N-periodic sequence, we will only be able to reconstruct the first N-F-states. Therefore, we need to shift the F-states information in k-space using a gradient spoiler as shown here, to make sure only N F-states are inside the acquired k-space.

## Slide 11
Theory
N-periodic ka-SPGR
F = 0
Correct for
phase modulation
F = 1
F = 2
F = 3
F = 4
Shift
F-states
(image-space)
F0
F1
F2
F3
F4
n = 0
n = 1
n = 2
n = 3
n = 4
Reconstruction example:
5-Periodic ka-SPGR
F0
F1
F2
F3
F4
F-state magnitude
Time
Notes: An example of reconstructing the F-states and fitting the T2* value is shown here using 5-periodic ka-SPGR. We correct the phase modulation to extract components and reconstruct F-states. Then we shift the information back to the middle of k-space and inverse Fourier transform to get the F-states signal. The magnitude of F-states is then used to fit exponential decay against the effective time and acquire T2* as shown here.
As we averaged across multiple acquisitions to reconstruct the T2* decay signal, this method is theoretically motion robust.
---------------------------------------------------------
An example of reconstructing the F-states and fitting the T2* value is shown here using 5-periodic ka-SPGR. We correct the phase modulation of each acquired signal using this equation and sum up to form the k-space of corresponding F-states. Then we shift the off-set F-states information back to the middle of k-space and inverse Fourier transform to get the F-states signal. The magnitude of F-states are then used to fit exponential decay against the effective time and acquire T2* as shown here.

## Slide 12
Theory
N-periodic ka-SPGR
F = 0
Correct for
phase modulation
F = 1
F = 2
F = 3
F = 4
Shift
F-states
(image-space)
F0
F1
F2
F3
F4
n = 0
n = 1
n = 2
n = 3
n = 4
Reconstruction example:
5-Periodic ka-SPGR
F0
F1
F2
F3
F4
F-state magnitude
Time
Quantitative analysis?
Scan parameter choice?
Notes: However, the accuracy of the T2* mapping of this technique is not quantitatively analysed and no optimal scanning parameters are suggested before.

## Slide 13
Method Overview
Monte Carlo computer simulation
Optimised TR and N for N-periodic ka-SPGR
Practical phantom scanning experiment
Percentage bias
Bias and variation
T2* measurement efficiency
T2* mapping image
Quantitative analysis?
Scan parameter choice?
Notes: So in this project, we focused on these two points. Here is an overview of the entire project, we first build a computer simulation to calculate bias and find optimal scanning parameters at the lowest bias. Then the optimized sequences are used to perform a practical phantom scanning experiment to generate T2* image and analyse the performance in reality compared with gold-standard method.

## Slide 14
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…

## Slide 15
Computer Simulation
Method
T2* ground truth (13-53 ms)
TR
(2-10ms)
1000 measured T2*
Monte Carlo experiment data
For 3-12 periodic ka-SPGR
Notes: So, for the computer simulation. We first model the formation of the signal with a commonly used simulation in MRI called Bloch simulation with the N-periodic ka-SPGR sequence setting here in figure (a) under the SN tissue environment. Considering the phase shift due to different off-resonance frequencies, we draw an off-resonance profile, and by doing the Fourier transform of this profile we get the simulated F-states. The field inhomogeneity effects caused by the iron can be calculated from known T2 and T2* and is modelled using a Lorentzian distribution.
Then we start modelling the acquisition by adding complex Gaussian noise to the F-states and a Monte Carlo experiment of 1000 measurements is carried out. We use only the first N F-states to perform exponential fit for measured T2*.
By the end of the simulation, we got simulated data for 3 periodic – 12 periodic, for TR between 2-10ms and ground truth T2* which is the PD biomarker value.
-------------------------------------------------------------------------
So, for the computer simulation. We first model the formation of the signal with a commonly used simulation in MRI called Bloch simulation, which is just a bunch of rotation and translation matrixes. We adjust the simulation according to the N-periodic ka-SPGR sequence setting here in figure a, with SN tissue properties T1 and T2 and optimal flip angle according to the Ernst equation. Considering the phase shift due to different off-resonance frequencies, we are able to draw an off-resonance profile, and by doing the Fourier transform of this profile we get the modelled F-states. The field inhomogeneity effects caused by the iron described by T2’ time constant can be calculated from known T2 and T2* and is modelled using a Lorentzian distribution.
Then we start modelling the acquisition by adding a complex Gaussian noise to the F-states and using only the first N F-states to perform exponential fit for T2*.
A Monte Carlo experiment of 1000 measurements is carried out. By the end of the simulation, we got simulated data for 3 periodic – 12 periodic, for TR between 2-10ms and ground truth T2* which is the PD biomarker value.

## Slide 16
Computer Simulation
Result
Average T2* percentage bias
7 or 12- periodic ka-SPGR
Notes: Averaged percentage bias is calculated for each TR and periodicity combination and plotted against TR. Due to the clinal scanner limit, we can only use TR greater than 6ms. The optimal point is found at minimum bias, which is 7/12-periodic with TR = 6ms.

## Slide 17
Computer Simulation
Result
Optimal ka-SPGR percentage bias and standard deviation
N = 12, TR = 6ms
N = 7, TR = 6ms
Notes: This shows the percentage bias and standard deviation of the 2 optimal sequences we selected. The middle green line is the zero bias line, with red dotted lines indicating +-5%. It shows they all inside +-5% bias line, and the 12-periodic had an almost half standard deviation than the 7-periodic.

## Slide 18
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…

## Slide 19
Practical MRI Scan
Method
Notes: The scan is performed on a phantom sphere, shown here, red error shows the layer of the sphere we scanned, the picture on the right is the cross-section we scanned and each of these small spheres has different T2* values.

## Slide 20
Practical MRI Scan
Method
N-periodic ka-SPGR
Time
SI
TE
TE + TR
TE + 2TR
TE + (N-1)TR
……
……
Notes: The two optimized sequences 7 and 12 periodic ka-SPGR with 6 ms TR are used to perform the scan.

## Slide 21
Practical MRI Scan
Method
Multi-echo SPGR
N-periodic ka-SPGR
Time
SI
TE1
TE2
TE3
TE
TE + TR
TEN
TE + 2TR
TE + (N-1)TR
……
……
……
Notes: As we want to compare N-periodic ka-SPGR with a gold-standard, multi-echo SPGR, we need to match contrast using the same image and voxel sizes. Also, we want the minimize the influence of the exponential fit algorithm, so we want to make sure they all acquire the same part of the decay. It is achieved by matching the multiple TEs with N-periodic effective time.

## Slide 22
Practical MRI Scan
Result
Gold-standard
Quantitative T2* mapping images
7-periodic ka-SPGR
12-periodic ka-SPGR
Notes: These are the quantitative T2* mapping images, the T2* value of fiducial spheres are colour coded on top of the MR greyscale image. Notice that the fiducial spheres with T2* value outside the PD detection range will be excluded for mapping and further analysis.
It clearly shows the N-periodic T2* mapping image aligns with the gold-standard map. And the healthy and PD SN T2*, 10 and 50ms be distinguished.

## Slide 23
Practical MRI Scan
Result
T2* percentage error
N = 7
N = 12
Effective T2* SNR
Notes: Then we further analysed the percentage error and effective T2* signal-to-noise ratio of each fiducial sphere.
For the percentage error, red lines indicate +-10%. We can see both the 7 and 12 periods stay within +-10% and the 12 period has a smaller standard deviation.
The effective T2* signal-to-noise ratio shows the efficiency of the technique. The orange bars show N-periodic ka-SPGR and the blue bars show the gold standard. It shows under the same contrast and environment, 7 periodic is significantly more efficient at low T2*, which indicates it might be a good detector for PD-related T2* fall.
--------------------------------------------------
Then we further analysed the percentage error and effective T2* signal-to-noise ratio of each fiducial sphere.
For the percentage error, the middle line is the zero bias, red lines indicate +-10%. We can see both the 7 and 12 periods stay within +-10% and the 12 period has a smaller standard deviation.
The effective T2* signal-to-noise ratio takes into account the effective scan time, it shows the efficiency of the technique. The orange bars show N-periodic ka-SPGR and the blue bars show the gold standard. It shows under the same contrast and environment, 7 periodic is significantly more efficient at low T2*, which indicates it might be a good detector for PD-related T2* fall.

## Slide 24
Development of MRI T2* Mapping Model for N-Periodic ka-SPGR Sequence:
A Promising Technique for Parkinson’s Disease Biomarker Detection
Background introduction
Theory
N-periodic ka-SPGR
Computer simulation
Method
Result
Practical MRI scan
Method
Result
Conclusion and Limitations
Q & A
Overview…

## Slide 25
Conclusion
Conclusion
Optimal N-periodic ka-SPGR:
7-periodic ka-SPGR TR = 6ms
12-periodic ka-SPGR TR = 6ms
12-periodic is more precise.
7-periodic is more efficient for PD detection.
Notes: From this project, we developed a computer simulation for N-periodic ka-SPGR sequence T2* mapping, from which, we obtained the optimal ka-SPGR sequence, which is 7/12 periodic ka-SPGR with 6 ms TR. The computer simulation result and phantom experiment result both suggest the 12-periodic is more precise, and 7-periodic is more efficient for PD detection.

## Slide 26
Limitation and Further
Limitation
Phantom misalign with substantia nigra tissue property (T1).
Large amounts of in vivo scan experiments are required to statistically prove the technique for clinical use.
Further research
The optimal parameters can be used to test motion robustness in vivo.
Notes: To prove the biggest advantage of this sequence.

## Slide 27
Q&A
Thanks for listening!
Any questions?
Notes: To prove the biggest advantage of this sequence.

## Slide 28: References:
Brammerloh, M., Morawski, M., Friedrich, I., Reinert, T., Lange, C., Pelicon, P., Vavpetič, P., Jankuhn, S., Jäger, C., Alkemade, A., Balesar, R., Pine, K., Gavriilidis, F., Trampel, R., Reimer, E., Arendt, T., Weiskopf, N., & Kirilina, E. (2021). Measuring the iron content of dopaminergic neurons in substantia nigra with MRI relaxometry. NeuroImage, 239. https://doi.org/10.1016/j.neuroimage.2021.118255
Holdsworth, S. J., Rahimi, M. S., Ni, W. W., Zaharchuk, G., & Moseley, M. E. (2016). Amplified magnetic resonance imaging (aMRI). Magnetic Resonance in Medicine, 75(6), 2245–2254. https://doi.org/10.1002/mrm.26142
Lally, P. J., Chiew, M., Matthews, P. M., Miller, K. L., & Bangerter, N. K. (n.d.). A motion-robust, short-TR alternative to multi-echo SPGR.
References